Standard Operating Procedure for PEALD Operation

Machine Appearance and Main screen

1

4

2

3

| 5 |  |
| --- | --- |
| 6 | 7 |

8

Machine appearance

- System control and operation area

- Load lock chamber

- Chamber A (High k)

- Chamber B (Metal)

Main screen

- Function menu

- Function submenu

- System diagram

- Load lock control options

Note before Use

1

2	3

4

1.Check if the machine status is “Normal” and make sure that there is no alarm message.

2.Set the temperature of “Sub Heater.” (TMA:250°C, TEMAZ:250°C, TEMAH:250°C)

3.Check if the temperature of Pump/Gas line is 130°C.

4.Check the temperature of “Precursor.” (Al2O3: 20°C, HfO2: 75°C, TEMAZ: 75°C)

Chip Placement and Transfer

3

4

2

1	5

- Tap on “Pump valve”→“Close.”

- Tap on “Vent valve”→“Open.” After making sure that the pressure is “ATM,” open the chamber and put chips in.

- Tap on Vent valve→Close and Pump valve→Open. After making sure that the pressure is “5E-3 Torr↓,” the chips can be transferred.

- Select the location to transfer the chips.

- Tap on “Auto Transfer” to start the transfer. After completion, make sure that there are no chips on the robot arm.

Process Parameter Editing

Screen1

Screen2

2

3

1

- Select the program name and make sure the name on the upper right is correct.

- Tap on “Edit Rec” to enter Screen 2.

- Tap on “Edit Step” to enter Screen 3.

- Follow the process steps for editing and leave this page after completion.

Screen3

4

Process Execution [1]

Screen2

4

5

Screen1

1

2

3

  - Tap on “Service” in the function menu

  - Tap on “Operation” in the function submenu to enter the “Chamber operation” screen.

  - Tap on “Operation” for the process chamber to enter Screen 2.

  - Tap on “Chamber Maintenance”→“On Line.”

  - Tap on “Previous” in the function submenu to return to Screen 1.

Process Execution [2]

Screen1

1

4

Screen2

3

2

- Select Chamber A: in the Chamber Process Program Selection block at the bottom of Screen 1, 	and a drop-down program list displays in Screen 2.

- Select the name of the required program.

- Make sure that the selected program is correct and tap on Recipe to confirm.

- Tap on “Recipe Run” to execute the program.

Real-time Process Monitoring

1

2

  - Tap on “Chamber A” in the function menu.

  - Tap on “Process” in the function submenu to enter the “Monitor chamber A process” screen to monitor process progress in real time.