C10-8-inch Chemical Wet Etching Workstation - Standard Process

| Process Conditions for the Cleaning Station |  |  |  |  |
| --- | --- | --- | --- | --- |
|  | SC-1 | SC-2 | DHF | SPM |
| Types of chemicals | NH4OH : H2O2 : H2O | HCl : H2O2 : H2O | HF : H2O | H2SO4 : H2O2 |
| Process concentration | 1 : 4 : 20 | 1 : 1 : 6 | 1 : 50 | 4 : 1 |
| Process temperature | 45°C | 45°C | 24°C | 125°C |
| Process time | 600sec | 600sec | 60sec | 600sec |
| Purpose of chemicals | Remove particles/organic pollutants | Remove metal pollutants | Remove native oxide layer | Remove organic pollutants |
| Process Conditions for the Photoresist/Wet Etching Station |  |  |  |  |
|  | WaferP.R. Strip (on non metal) | H3PO4 | B.O.E. 7:1 |  |
| Types of chemicals | H2SO4 : H2O2 | H3PO 4 | B.O.E. 7:1 (NH4F:HF) |  |
| Process concentration | Pure H2SO4 + 100 ml H2O2 | H3PO 4+ H2O 200 ml | Pure B.O.E. |  |
| Process temperature | 125°C | 150°C | Ambient temperature |  |
| Process time | 600sec | Depends on required etching thickness | Depends on required etching thickness |  |
| Purpose of chemicals | Remove organic pollutants | Remove Si3N4 | Etch silicon dioxide |  |

Process conditions currently open

Cleaning Station

- Before the long gate oxide layer/Before the manufacturing of structures directly related to electrical properties/Growth of films with a thicknesses of less than 100Å

RCA flower: SPM → QDR → DHF → QDR → SC-1 → QDR → SC-2 → QDR → DHF → QDR → SPIN DRY

- Cleaning before general high-temperature and high vacuum processes/forming film thicknesses less than 100Å

STD flower: SC-1 → QDR → SC-2 → QDR → SPIN DRY

- Cleaning before general high-temperature and high vacuum processes, but worried about the effect of native oxide layer

STD + DHF dip flower: SC-1 → QDR → SC-2 → QDR → DHF → QDR → SPIN DRY

- After etching/cleaning after non-metal chemical mechanical polishing/removing side wall polymer

SC-1 → QDR → SPIN DRY

- Only when removing native oxide layer

DHF → QDR → SPIN DRY

- Cleaning before the Poly Silicon TFT process gate oxide

SPM → QDR → DHF → QDR → SC-2 → QDR → DHF → QDR → SPIN DRY

Photoresist/Wet Etching Station

- Removing front-end process photoresist

SPM#2 → QDR → SPIN DRY

- SiO2/TEOS wet etching (entry of wafers with photoresist patterns is permitted)

B.O.E. 7:1 → QDR → SPIN DRY

- Removal of the front-end process silicon nitride (Si3N4) film

H3PO4 → QDR → SPIN DRY