8” fully automated surface profiler

- It is forbidden to press hard down the holder and stage by hand.

- Before logging in P10, don’t touch the keyboard and track ball.

- Scan depth: 300Å~300µm (Resolution 1 or 25 Å)

- Scan Length:

- Scan length must not be greater than 60mm

- Scan time must not exceed 100sec

- Press the Space bar to change the unit within a given unit system (ex. µm to mm; min. to mils)

- Stylus Force:

- Recommended Stylus Force Versus Scan Speed

| Speed (µm/s) | Recommended Force (mg) | Speed (µm/s) | Recommended Force (mg) |
| --- | --- | --- | --- |
| 1 | 1-25 | 100 | 10-35 |
| 2 | 1-25 | 400 | 15-44 |
| 5 | 3-25 | 1000 | 20-50 |
| 10 | 5-25 | 2000 | 25-60 |
| 20 | 7-25 | 5000 | 30-75 |
| 50 | 8-30 | 10000 | 35-85 |
| - | - | 25000 | 40-100 |

- The force decreases approximately 1mg every 13um below the center. Therefore, use at least 20mg when in the 300um range.

- Contact Speed (1-10):

- This parameter specifies the stylus drop speed at contact with the substrate.

- Soft substrates such as aluminum or photoresist can require a setting of 1or 2. We recommend slower contact speed settings of 1 to 3 with our stylus.

- Vertical Range/Resolution:

- Select 13µm for the feature heights on the order of 6.5µm

- Select 300µm for the feature heights on the order of 150µm