| Document Title: | Machine SOP (CF-E14 8” fully automated surface profiler) |
| --- | --- |
| Document No.: | Q3-NL04 |
| Draft Dept.: | Film etching division |
| Draft Date: | 2021-10-06 |

| Version | Edited by | Effective date | Approve document No. | Revision/change description | Revised Page |
| --- | --- | --- | --- | --- | --- |
| 1.0 | Po-Hao Chu |  |  | New release | --- |

# Purpose:

Define 8” fully automated surface profiler practices to ensure production quality

# Scope:

Applicable to 8” fully automated surface profiler.

# Rights and Responsibilities:

- Organization responsibilities: The process engineer is responsible for the drafting and revision of specifications.

- Equipment manager is responsible for handling anomalies to maintain normal operation.

- Qualified operator: Operators passing Tencor P-10 surface profiler operation test.

# Definition of terms:

No.

# Relevant Documents:

8” fully automated surface profiler user manual.

# Standard Operating Procedure:

- Inspect machine condition and current status (green card: in use; red card: under maintenance; yellow card: under test). Boot the machine on the web (web on) and record the start time in machine usage log.

FOCUS

MAN LOAD

- Move out the carrier by clicking MAN LOAD.

- Open the clear door, put the wafer on the carrier (flat side outward) and turn on vacuum suction.

- Close the clear door, click MAN LOAD to move the carrier to the position for measurement (DO NOT open the clear door during operation).

- Click FOCUS to move the stylus to the sample.

- Use mouse buttons to select the scan area or the trackball to control X-Y.

- After determining the scan area, click “Cancel” to access Recipe screen and select the required recipe.

- Check or change the scanning program parameters and click “START" to start scanning. After scanning, the system will automatically jump to the screen below (Fig. 1).

Cancel

START

Fig. 1

- If the result is skew, click Level to call L-L and R-R to level the scanned profile.

- Press 2 keys to translate and 1 key to adjust the width. Adjust L-L and R-R to the same level and leave some space to acquire the average.

- Click Level again to level the image.

- Place L-L and R-R at the height (or width) for comparison and the height (or width) difference will be displayed on the left side of screen.

- After measurement, move out the carrier by clicking MAN LOAD.

- Open the clear door, turn off vacuum suction and take out the wafer.

- Turn off the machine on the web (web off) and record the end time in machine usage log.

# Reference forms and attachments:

- Q4-NL02 Machine management card

- Q4-NL03 Machine evaluation checklist

- Q4-NL04 Machine checklist

- Q4-NL06 Anomaly and Preventive/ correction action list