Standard Operating Procedures for Focused Electron Beam System

- Placement of the specimen:

  - Turn off E beam/I beam,

  - Click unload loadlock

  - Open loadlock door and place the sample in, and click Load

Completed

Not completed

Check if the vacuum degree of the chamber is below 5x10-5 mbar

- ON/OFF of electron and ion beams

Click Wake Up to turn on E beam/I beam

- Search for specimen/cutting point

  - Focus on and search for the sample with low-magnification SEM (mode 1)

  - Perform Z to FWD connection,

  - Set Z=4.25 mm (in SEM mode)

  - Adjust magnification to approx. 5000x (the scale bar showing 5um), and adjust the focus (Note: the stage must be moved with the mouse and the QUADE 1-3 in the navigator screen)

- Setting the eucentric height of the specimen (SEM mode)

  - The Z-axis must be moved with the mouse and the navigator screen

Z to FWD connection and set Z=5mm

SEM focus is clear (SEM mode 1)

Adjust Z to Y offset<5um

Check if Y offset<5um

  - Parameter settings of the coordinates of 5 axes

- Cutting of the specimen and coating with a Pt protective layer

  - Switch to FIB mode (select 30~50pA for beam current)

  - Select Patterning

  - Select patterns, set xyz, and select material file.

  - Select appropriate I beam currents (current density must be lower than 10pA/um2)

  - Heat Pt and insert the needle

  - Run patterning

  - Run X-section

Select regular cross section-> cleaning cross section

- Capturing and saving images (switching to SEM mode)

| Real-time image |  |
| --- | --- |
| Capture single image |  |

or press F2 for the default photo settings