| Machine Name: | N&K1500 Film Thickness Measurement System |
| --- | --- |
| Machine No.: | CF-M23 |
| Prepared by: | Lithography and Mask Division |
| Issue Date: | Sep. 8, 2021 |

Document Issue and Revision History

| NO. | Issue Date | Revision Application No. | Issued/Revised Content | Revised Page | Version |
| --- | --- | --- | --- | --- | --- |
| 01 | Nov. 1, 2013 | 102-DC21 | First issue | ----- | 01 |
| 02 | Nov. 1, 2021 |  | Issue of ver. 1.1 | ----- | 1.1 |
| 03 |  |  |  |  |  |
| 04 |  |  |  |  |  |
| 05 |  |  |  |  |  |
| 06 |  |  |  |  |  |
| 07 |  |  |  |  |  |
| 08 |  |  |  |  |  |

| Approved by | Reviewed by | Prepared by |
| --- | --- | --- |

- Purpose:

Defining the operational standards of the N&K1500 Film Thickness Measurement System to ensure the quality of its operation.

- Scope:

Applicable to the N&K1500 Film Thickness Measurement System.

- Responsibilities:

  - Organizational responsibilities: The engineer is responsible for the drafting and revision of standards.

  - Personnel qualified for execution: Personnel who have passed the test for the N&K1500 Film Thickness Measurement System.

- Definition of Terms:

None.

- Related Documents:

None.

- Standard Operating Procedures:

  - Enter MES to swipe card to start up the machine.

  - Place the 8-inch baseline wafer into the stage.

  - Turn on vacuum.

  - In Library, select Size as 8 inch (stage size) and click associate.

  - Check if the option of 8” is ticked under Mapping.

  - Execute Baseline.

  - The baseline total value must be less than or equal to 0.20% to be acceptable.

  - Remove the baseline wafer and place the wafer to be measured.

  - Use mapping to select the location to be measured.

  - To enter the film structure, use the RD method or the built-in program in Library.

  - Data collection and analysis.

  - Select Single Measurement or Multiple Point Measurement, and click START.

  - After keying in data information, click “ok”.

  - Click NO, as there is no need to save the file.

  - The goodness of fit after measurement must be at least 0.98 for the measurement data to be accurate.

  - Assuming multiple points are measured, click the icon for statistics.

  - You can check the data of multiple points.

- Applicable Forms and Attachments:

  - Equipment Evaluation Checklist S4-NL01A

  - Equipment Inspection Checklist S4-NL01B

  - Equipment Use Log Sheet S4-NL01C

  - Form for Irregularities and Their Correction and Prevention S4-NL02A

  - Instrument and Equipment Management Card S4-NL04B