Introduction of S9260A In-line Electron Microscope

| Instrument no.: CF-L16 |  |
| --- | --- |
| Instrument name: S9260A In-line Electron Microscope |  |
| Storage location: Class 10 |  |
| Model: HITACHI S920A CD SEM |  |

Specifications of the Machine

| Item | Specifications |
| --- | --- |
| Specimen size | 8” wafer (SEMI or JETDA standard) |
| Accelerating voltage | 500-1600 V |
| Resolution | 3nm (800V) |
| Magnification | 500X to 300K X (SEM). Guarantee: 4000 X-200K X 110 X (OM) |
| Measurable line width | 0.1-2 um |
| Reliability | - 1% or 3 nm (3 sigma) |
| Moving range (X or Y) | 0-200 mm |
| Transfer mechanism | Automatic transfer 3 SMIF port (8” only) |
| Monitor | EWS 21 monitor |
| Capacity | Max. 50 specimens/hr (5 measuring points/specimen) |
| Storage | Built-in hard drive 3.5 floppy |

Functions of the Machine

The S9260A In-line Electron Microscope is mainly used for measuring the line width of in-line products. It is characterized by the fact that it does not require a wafer to go through the pre-processing steps such as cutting or metal coating to observe and measure the patterns of photoresist, insulation layer and metal layer. This measuring instrument uses non-destructive testing, so that any wafer whose measurement is completed may proceed to the next step in the process.

The machine can only measure line width. Since a specimen cannot be tilted, it is not possible to observe the specimen section.

Restrictions on the Machine (Materials, Range of Parameters, etc.)

Notes:

- The specimen must be of the required size and materials: 8 Si specimen.

- The specimen must be complete, with no missing corners or cracks.

- Due to the high vacuum in the electron gun, the specimen must not be any volatile substance or volatile during electron beam scanning.

- If the specimen contains any photoresist, it must undergo a hard bake before it is fed into the gun.

- The specimen must not contain any magnetic material.